## Slide 1
Richard N. Walters
Richard E. Lyon & Natallia Safronava
Federal Aviation Administration
William J. Hughes Technical Center
Fire Safety Branch ANG-E21
Atlantic City, New Jersey 08405
Preliminary Results from the MCC Inter-Laboratory Study
Federal Aviation
Administration

## Slide 2
FAA Microscale Combustion Calorimeter
(ASTM D7309-21)
The MCC test measures Materials Properties related to Flammability on a milligram size scale.
One of the outputs: Fire Growth Capacity (FGC).
The FGC is a measure of ignitability and burning rate of the material, i.e., the total fire hazard
MCC is proposed method for alternate means of compliance when a small change is made to a construction

## Slide 3
FAA Microscale Combustion Calorimeter
HRR by
O2 Consumption

## Slide 4
FAA Microscale Combustion Calorimeter
HRR by
O2 Consumption

## Slide 5: Standard Test ASTM D 7309 (Method A)
Total Heat Release
Heat of Combustion, Q
(Fire Load)
0
100
200
300
400
500
600
700
450
500
550
600
650
700
Temperature (C)
Onset (Ignition)
Temperature
Peak Pyrolysis
Temperature
Peak HRR &
Heat Release Capacity, c
Pyrolysis Residue 
(Char Fraction)
Heat Release Rate Q (W/g)
Anaerobic pyrolysis at 1 K/s + complete combustion of gases at 900C, 20% O2
Microscale Combustion
Calorimeter

## Slide 6
MCC procedure for FGC
0
10
20
30
40
50
60
70
80
0
2
4
6
8
10
200
300
400
500
600
700
800
900
Specific Heat Release Rate, Q (W/g)
Temperature, T (°C)
Q (J/g)
Q∞
T1
T2
-5% Q∞
+5%
Q∞
T0 = 25°C (298K)
T1 = Ignition temperature
T2 = Burning temperature
Measure specific heat release rate Q versus temperature T as per ASTM D7309 (5 replicates)
Integrate Q/ versus T to obtain Q versus T, i.e., Q(T )
Obtain total heat release
Q(T∞) = Q∞ = hc(J/g)
Obtain T1 at 5% deflection from Q(T) baseline, i.e., at 0.05Q∞
Obtain T2 at Q∞ i.e., 0.95Q∞.
Calculate Fire Growth capacity (FGC)
Fire Growth Capacity - FGC

## Slide 7
Commodity Plastics
Engineering Plastics
Specialty Plastics
Heat Resistant Plastics
Fire Resistant Polymers
Fire Growth Capacity, FGC (J/g-K)

## Slide 8
Small changes in the composition of certified aircraft cabin materials are often needed due to unavailability of the original components or environmental regulations
Recertification of the entire constructions are costly
Aircraft manufacturers and suppliers asked the FAA to explore alternative means of complying with the FAR in 2015.
The Material Similarity Task Group was created to develop a method and criterion for comparing flammability of samples at micro-scale using ASTM D7309 (MCC).
Similarity Project Background

## Slide 9
Similarity Criterion
The no-effect level of a material change at micro-scale is equal to the relative uncertainty of bench-scale FAR fire test results at the 95% confidence level, i.e., 30%, as determined by testing.
__________________________________________
Next steps:
Jeff Gardlin’s presentation on Transport Airplane Issue List (TAIL).
ASTM D7309-21 inter-laboratory study to-
  Determine repeatability and reproducibility of FGC
  Include ultra-low heat release aircraft phenolic resin.

## Slide 10
Similarity Criterion
The no-effect level of a material change at micro-scale is equal to the relative uncertainty of bench-scale FAR fire test results at the 95% confidence level, i.e., 30%, as determined by testing.
__________________________________________
Next steps:
Jeff Gardlin’s presentation on Transport Airplane Issue List (TAIL).
ASTM D7309-21 inter-laboratory study to-
  Determine repeatability and reproducibility of FGC
  Include ultra-low heat release aircraft phenolic resin.

## Slide 11: ASTM E691
Conducting an Inter-Laboratory Study to Determine the Precision of a Method
Repeatability
  An action, event, or other thing that is done again
Reproducibility
  Create something very similar to (something else) in a different medium or context

## Slide 12: Measured Values
Sample Weight (mg)
Peak Heat Release Rate (W/g)
Peak Heat Release Temperature (C)
Total Heat Released (kJ/g)
Fire Growth Capacity (J/g-K)
Char Yield (%)

## Slide 13: Inter-Laboratory Study – Round 1
4 manufacturers/licensees of the MCC
  Deatak
  Fire Testing Technologies
  Concept Equipment Ltd
  Me
Polystyrene sample sent to labs
  Over 20 labs were asked to participate
  Several labs could not participate due to equipment not working
Data received from 14 labs
  Several labs need equipment adjusted
  Several labs equipment was fine but data was re-analyzed and interpreted by me

## Slide 14: Lab Comparison – Peak Heat Release Rate
Chart: Lab
Deatak P | 1184.3333
Deatak C | 1189
UMASS C | 1026.6667
SGS C | 1090.4667
UDRI C | 1149
BAM | 1070.6667
NIST | 1141
FAA C | 1124.6667
FAA P | 1102
3M | 1108
Boeing 1923 | 1187.0333
Boeing 1924 | 1187.1
1118 W/g
1 2 3 11 12 14
±5.6%
±63 W/g

## Slide 15: Lab Comparison – Total Heat Release
Chart: Lab
Deatak P | 41.5333
Deatak C | 41.6
UMASS C | 38.1333
SGS C | 41.7
UDRI C | 41.0667
BAM | 39.8333
NIST | 40.2183
FAA C | 41.4667
FAA P | 40.8667
3M | 40.2333
Boeing 1923 | 40.7
Boeing 1924 | 40.2667
3 4 14
40.6 kJ/g
±2.8%
±1.1 kJ/g

## Slide 16: Lab Comparison – Fire Growth Capacity
Chart: Lab
Deatak P | 899.6667
Deatak C | 908.3333
UMASS C | 810.8
SGS C | 889.9667
UDRI C | 877.3333
BAM | 854.985
NIST | 889.5333
FAA C | 846.6667
FAA P | 847.6667
3M | 845.3333
Boeing 1923 | 887
Boeing 1924 | 863
2 3
868 J/g-K
±4.1%
±35 J/g-K

## Slide 17: Lab Comparison – Peak Temperature
Chart: Lab
Deatak P | 446.0307
Deatak C | 447.3667
UMASS C | 440
SGS C | 439.6667
UDRI C | 450.2667
BAM | 441.3333
NIST | 448.7333
FAA C | 440.3333
FAA P | 435.7297
3M | 446.3333
Boeing 1923 | 433
Boeing 1924 | 443.3333
5 7 11 14
441 C
±1.5%
±6 C

## Slide 18: Sources of Error
Heat Release Rate
  Bad O2 sensor/calibration
  Leak in system
  Sample size
  Bad weight measurements
  Flow controllers/meter out of calibration
Temperature
  Thermal lag
  DAQ board calibration
  Time shift

## Slide 19: Future Work
Offer suggestions to operators to fix errors
Inter-laboratory study – Round 2 (ASTM)
Send set of 5-6 samples to labs
Results submitted directly to ASTM
Updated Precision & Bias statement
FAA report
Chart: 
Chart: Aircraft Phenolic

## Slide 20: FAA Reports - https://www.fire.tc.faa.gov/
ILS
MCC
FGC
Similarity

## Slide 21: Contact Info
Richard Walters, PhD
Federal Aviation Administration
William J Hughes Technical Center
Fire Safety Branch ANG-E21 Bldg 277
Atlantic City International Airport, NJ 08405
609-515-9508
richard.walters@faa.gov