## Slide 1: Inter-laboratory Study of Radiant Panel Test for Measuring Flammability of Thermal/Acoustic Insulation Materials

## Slide 2: Introduction
New Handbook chapter being written
Need more data to finalize changes
Industry dimension survey
Round Robin
Air Flow Study
  Varied openings around the sliding platform to test how it affected air flow and material test results.
  Future Work

## Slide 3: Dimension Survey

## Slide 4: Dimension Survey

## Slide 5: Dimension Survey
Gaps around the drawer are the biggest difference between machines
Chart: Open Area Around Drawer
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | Y | Z | A2
 | 123.75 | 0 | 162.75 | 216.3281 | 201.5 | 0 | 116.375 | 66.9475 | 197.2372 | 120.25 | 79.875 | 124.8125 | 123.7656 | 183.7 | 279 | 183.92 | 13.3 | 262.875 | 112.665 | 308.1875 | 177.7656 | 0 | 0 | 6.25 | 58.1 | 188.9812 | 235.7188

## Slide 6: Round Robin
Materials sent out to 28 labs (including FAA)
24 Responses Received
12 samples for each lab

## Slide 7
Chart: Metalized PEEK
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | Y | Z | A2 | B2
Flame Propagation | 0.2 | 0.2 | 0.2 | 0.2 | 1.5 | 1 | 0.2 | 2.3 | 3.74 | 4.01 | 2.5 | 2.75 | 1 | 1.7 | 1.875 | 2.875 | 1.5 | 1.375 | 0.7 | 0.7 | 0.9 | 1 | 6.4 | 5.7 | 1.37 | 2.7 | 0.7065 | 2.7925 | 2 | 1.5 | 5.5 | 5.5 | 1 | 0.5 | 5.04 | 4.37 | 1 | 1.5 | 1.2 | 1.22 | 1.5 | 1.5 | 1 | 3 | 0.93 | 1.01 | 4.5 | 5.5
After Flame | 0 | 0 | 0 | 0 | 15 | 0 | 0 | 8 | 2.6 | 13.8 | 2.38 | 2.89 | 0 | 0 | 0 | 0 | 3.18 | 0 | 0 | 0 | 0 | 0 | 10.1 | 16.6 | 0 | 4.7 | 2.2 | 10 | 9.4 | 0 | 5.8 | 3.9 | 0 | 0 | 8.337 | 12 | 0 | 0 | 0 | 0 | 2.3 | 3.8 | 0 | 5.4 | 0 | 0 | 12 | 17.4
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 8
Chart: Unmetalized PEEK
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 9
Chart: Metalized PEEK w/ Tape
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 10
Chart: Polyimide
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | Y | Z | A2 | B2
Flame Propagation | 0.2 | 0.2 | 0.2 | 0.5 | 0.2 | 0.2 | 0.2 | 0.2 | 1.57 | 1.1 | 0.5 | 0.5 | 0.7 | 0.5 | 1.25 | 1.125 | 0.625 | 0.625 | 0.6 | 0.5 | 0.6 | 0.9 | 0.9 | 0.9 | 1 | 1 | 0.994 | 1.0175 | 0.5 | 1 | 1 | 1 | 0.5 | 0.5 | 0.71 | 1.1 | 0.5 | 0.75 | 1.42 | 1.02 | 0.5 | 0.6 | 1 | 1 | 0.92 | 1.08 | 0.25 | 0.25
After Flame | 1.45 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.3 | 1.1 | 9.4 | 1.9 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.7 | 0.5
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 11: Polyimide
1-12
13-24

## Slide 12: Polyimide
Same test, completely different results.
FAA’s test result

## Slide 13: Air Flow Study
New handbook has a larger drawer length compared to the current handbook.
This was done to allow less air to flow into the chamber during testing.
Larger drawer isn’t necessary – the openings around the drawer are what we need to standardize.
We ran some tests to see how changing the size of these openings affects calibration and test results.

## Slide 14: Air Flow Study
Normal openings around the FAA’s radiant panel
Left | Right | Front | Back
5/16” | 1/8” | 0” | 2.25”

## Slide 15: Air Flow Study
Closed off left, rear, and right gaps

## Slide 16: Air Flow Study
2.125” Left Gap
1.875” Right Gap
Gaps around the drawer fully open

## Slide 17: Air Flow Study
 | Closed Around Drawer | NormalAround Drawer | Open w/ ~2” Gap on Each Side
Open Area Around Drawer(in2) | 2.09 | 123.77 | 183.44
Panel Set Point (°F) | 1096 | 1105 | 1137
Heat Flux (BTU/ft2s) | 1.501 | 1.502 | 1.495
Chamber Temperature(°F) | 381 | 366 | 425
Chimney Air Velocity (FPM) | 205 | 260 | 278
Chart: Calibration vs. Drawer Openings
| 2.09375 | 123.765625 | 183.4375
Panel Set Point | 1096 | 1105 | 1137
Chimney Air Velocity | 205 | 260 | 278

## Slide 18: Air Flow Study
Set Point (°F) | 1095 | 1110 | 1125 | 1137
Heat Flux (BTU/ft2s) | 1.398 | 1.439 | 1.481 | 1.495
Thermocouple (°F) | 385 | 395 | 410 | 425
Chimney (FPM) | 245 | 255 | 268 | 278
Under Right Side of Drawer (FPM) | 90 | 91 | 95 | 95
Under Left Side of Drawer (FPM) | 65 | 60 | 65 | 65
Chart: Air Flow vs. Set Point
| 1095 | 1110 | 1125 | 1137
Heat Flux | 1.398 | 1.439 | 1.481 | 1.495
Chimney Air Velocity | 245 | 255 | 268 | 278
Openings Around Drawer: 2.125” left, 1.875” right, 2.25 rear
As the panel temperature increases, the air flow into the chamber increases causing you to need to increase the panel temperature further to compensate

## Slide 19: Round Robin Data
Chart: Set Point vs Open Area Around Drawer
| 0 | 162.75 | 216.328125 | 201.5 | 0 | 66.947500000000005 | 197.23719999999997 | 120.25 | 79.875 | 123.765625 | 183.7 | 279 | 183.92 | 13.3 | 262.875 | 112.66499999999999 | 308.1875 | 177.765625 | 0 | 0 | 6.25 | 58.099999999999994 | 188.98124999999999 | 235.71875
 | 1040 | 1090 | 0.72 | 1163 | 1105 | 1091 | 1139 | 1055 | 868 | 1100 | 1113 | 1120 | 1117 | 1095 | 995 | 1141 | 1125 | 1079 | 1133 | 1123 | 860 | 1089 | 1079 | 1130

## Slide 20: Round Robin Data
Chart: Open Area vs. Chimney Velocity
| 0 | 162.75 | 216.328125 | 201.5 | 0 | 66.947500000000005 | 197.23719999999997 | 120.25 | 79.875 | 123.765625 | 279 | 183.92 | 13.3 | 112.66499999999999 | 308.1875 | 177.765625 | 0 | 0 | 6.25 | 58.099999999999994 | 188.98124999999999 | 235.71875
 | 0 | 254 | 0 | 329 | 231 | 203 | 256 | 215 | 200 | 255 | 208 | 0 | 38 | 217 | 314 | 181 | 190 | 338 | 0 | 110 | 256 | 285
Chart: Panel Set Point vs. Chimney Velocity
| 1040 | 1090 | 0.72 | 1163 | 1105 | 1091 | 1139 | 1055 | 868 | 1100 | 1120 | 1117 | 1095 | 1141 | 1125 | 1079 | 1133 | 1123 | 860 | 1089 | 1079 | 1130
 | 0 | 254 | 0 | 329 | 231 | 203 | 256 | 215 | 200 | 255 | 208 | 0 | 38 | 217 | 314 | 181 | 190 | 338 | 0 | 110 | 256 | 285

## Slide 21: Air Flow Study
Chart: Gaps Closed - Metalized PEEK
Flame Prop | 2.3 | 1 | 2.6 | 2.4 | 2.4
After Flame | 4 | 0 | 11 | 10 | 8.5
FP Fail | 2 | 2
AF Fail | 3 | 3
Chart: Gaps Open - Metalized PEEK
Flame Prop | 1 | 2.3 | 1.1 | 0.9 | 0.9
After Flame | 0 | 0 | 0 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 22: Air Flow Study
Chart: Gaps Closed - Unmetalized PEEK
Flame Prop | 0.9 | 1.7 | 0.9 | 1.3 | 1.1
After Flame | 0 | 0 | 0 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3
Chart: Gaps Open - Unmetalized PEEK
Flame Prop | 1 | 1.1 | 1 | 1.1 | 1.2
After Flame | 0 | 0 | 0 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 23: Air Flow Study
Chart: Gaps Closed - Shrinking Material
Flame Prop | 0.7 | 0.7 | 0.7 | 1.6 | 0.8 | 0.9
After Flame | 0 | 0 | 0 | 2.8 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3
Chart: Gaps Open - Shrinking Material
Flame Prop | 0.5 | 0.7 | 0.6 | 1.2 | 0.8 | 0.7
After Flame | 0 | 0 | 0 | 0 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 24: Air Flow Study
Chart: Gaps Closed - Metalized PEEK with Tape
Flame Prop | 0.9 | 0.6 | 1.1 | 0.7 | 0.9
After Flame | 1 | 0.5 | 2 | 0 | 0.2
FP Fail | 2 | 2
AF Fail | 3 | 3
Chart: Gaps Open - Metalized PEEK with Tape
Flame Prop | 0.9 | 0.8 | 0.6 | 0.5 | 0.6
After Flame | 0.2 | 0 | 0 | 0 | 0
FP Fail | 2 | 2
AF Fail | 3 | 3

## Slide 25: Air Flow Study
At the June meeting, we finalized plans to further study this issue
Continue conducting tests with 3 different air gap levels
  Fully open
  Partially open
  Fully closed
Place array of thermocouples in the retaining frame to test how material temperature changes
Place smoke bomb below chamber to observe how the air travels through it
Additional material tests

## Slide 26: Air Flow Study
Array of 15 thermocouples placed inside retaining frame
Tested at each air-gap configuration
Calibrated with calorimeter to 1.5 Btu/ft2s each time
3 tests with each configuration
Temperature averaged over 5 minutes

## Slide 27: Air Flow Study

## Slide 28: Air Flow Study

## Slide 29: Air Flow Study

## Slide 30
Intermediate Gaps
Right: 0.125”
Left: 0.3125”
Rear: 2.25”
Front: 0”
Set Point: 1113°F
Chamber Temp: 370°F
Fully Closed
Right: 0”
Left: 0”
Rear: 0”
Front: 0”
Set Point: 1112°F
Chamber Temp: 384°F
Fully Open
Right: 1.875”
Left: 2.125”
Rear: 2.25”
Front: 0”
Set Point: 1158°F
Chamber Temp: 467°F

## Slide 31: Air Flow Study
Right Closed
Right: 0.125”
Left: 2.125”
Rear: 2.25”
Front: 0”
Set Point: 1113°F
Chamber Temp: 369°F
Right Open
Right: 1.875”
Left: 0”
Rear: 0”
Front: 0”
Set Point: 1126°F
Chamber Temp: 360°F

## Slide 32: Future Work
Send thermocouple array to 3 other labs to test
Statistically determined that at least 20 samples of a material would be required at each air gap level to get the best results
60 Metalized PEEK samples total for each lab

## Slide 33: Questions?